## Slide 1
Brian Magi
Associate Professor of Atmospheric Sciences
UNC Charlotte (Charlotte, North Carolina)
Calvin Cupini
Citizen Science Program Manager
Clean Air Carolina (Charlotte, North Carolina)
Comparison of PM2.5 measured in urban North Carolina settings using a low-cost optical particle counter and Federal Equivalent Methods
Presentation at CMAS Conference
23 October 2019

## Slide 2
Collaborations and input from
Mecklenburg County AQ (Charlotte, North Carolina)
Forsyth County AQ (Winston-Salem, North Carolina)
EPA National Exposure Research Lab
EPA National Health and Environmental Effects Lab
Students at UNC Charlotte

## Slide 3
Showing ~50 of the nearly 100 low-cost monitors in North Carolina
North Carolina has ~20 PM2.5 regulatory monitors total
PM2.5 () reported by low-cost monitors
Charlotte
Raleigh-Durham
Winston-Salem
Greensboro
Notes: Snapshot at 0830 Eastern, 13 October 2019

## Slide 4
Notes: Snapshot at 0830 Eastern, 13 October 2019 during Caples Escaped Rx Fire south of Lake Tahoe https://wildfiretoday.com/2019/10/12/caples-fire-spreads-to-the-south/ and image from https://wildfiretoday.com/2019/10/11/prescribed-fire-escapes-on-eldorado-national-forest/

## Slide 5
Notes: 2.5 million in Sacramento Metro, 5 million in SF Metro

## Slide 6
Good description at https://www.aqmd.gov/aq-spec/product/purpleair-pa-ii but also at Zamora et al 2019 Environmental Science & Technology https://doi.org/10.1021/acs.est.8b05174
Measures number concentration using 2 Plantower PMS5003 sensors. Reports number and mass concentrations.
Also measures temperature and relative humidity
Purple Air PA-II sensor package
Notes: Describe what we learn from the data, rather than how we learning from the data

## Slide 7
Data Objectives
What is the quality of the data from Purple Air PA-II Plantower sensors?
How long does a PA-II report consistent data?
CMAS Question
How can we make this data useful to modelers?

## Slide 8
CMAS
Co-location study sites

## Slide 9
Co-location
PA-II and FEM BAM 1022
Since March 2017
Co-location
PA-II and FEM T640X
Since March 2018
Co-location
PA-II and FEM BAM 1020
Since June 2019
Notes: Double collocation at Garinger HS = 4 sensors, at Remount = 3 (one is kaput), and single at Hatties Ave = 2 for a total of 9 data streams extending over months to as much as 2.5 years (Remount)

## Slide 10
Field Work in North Carolina

## Slide 11
Histograms of quality controlled hourly-averaged PM2.5 reported at 3 sites
Notes: This is only one of the two sensors (A vs B) and does not show Remount2

## Slide 12
Quality control and data processing
Collect raw sub-hourly data
QC Step 1 (non-physical reporting)
QC Step 2 (repeating artifact)
Hourly average the sub-hourly
Discard values less than limit of detection
Discard values greater than limit of sampling
Use remaining “uncorrected” PA-II data as input to a linear regression
of raw data discarded
of remaining data discarded
Notes: Example of Filters 5 and 6 using the data in this discussion (5/19/2019)
18,000+ hours or 760 days of samples before filtering.
14,000+ hours or 600 days of samples after filtering = 22% of the data is discarded.

## Slide 13
Correction of PA-II PM2.5 data using a linear regression
bias correction
RH correction
T correction
bias correction

## Slide 14
MLR
With , , ,
RMS uncorrected
Bias uncorrected
Charlotte co-location (17701 hours = 737 days)
RMS corrected for bias
RMS corrected for bias+RH
Notes: Bias correction , RH correction
Bias correction 1.24, RH correction 0.105

## Slide 15
MLR
With , , ,
RMS uncorrected
Bias uncorrected
Winston-Salem co-location (8432 hours = 355 days)
RMS corrected for bias
RMS corrected for bias+RH
Notes: 1.84, 0.111

## Slide 16
Site Sensor A | N (hours) | Bias before () | RMS before () | RMS after ()
Remount1 | 13964 | 5.3 | 7.8 | 4.2
Remount2 | 2892 | 9.8 | 12 | 4.5
Forsyth | 8432 | 7.3 | 9.6 | 2.9
Garinger1 | 1593 | 8.1 | 10 | 3.1
Garinger2 | 1435 | 7.1 | 9.5 | 3.2
Site Sensor A | N (hours) |  |  | 
Remount1 | 13964 | 5.3 | 7.8 | 4.2
Remount2 | 2892 | 9.8 | 12 | 4.5
Forsyth | 8432 | 7.3 | 9.6 | 2.9
Garinger1 | 1593 | 8.1 | 10 | 3.1
Garinger2 | 1435 | 7.1 | 9.5 | 3.2
Site Sensor B | N (hours) | Bias before () | RMS before () | RMS after ()
Remount1 | 17701 | 6.5 | 9 | 4.1
Remount2 | 2882 | 8.2 | 11 | 3.9
Forsyth | 8312 | 6.8 | 8.9 | 2.6
Garinger1 | 1611 | 7.8 | 9.9 | 3.1
Garinger2 | 1446 | 7.8 | 9.9 | 3.2
Site Sensor B | N (hours) |  |  | 
Remount1 | 17701 | 6.5 | 9 | 4.1
Remount2 | 2882 | 8.2 | 11 | 3.9
Forsyth | 8312 | 6.8 | 8.9 | 2.6
Garinger1 | 1611 | 7.8 | 9.9 | 3.1
Garinger2 | 1446 | 7.8 | 9.9 | 3.2
Comparison of all FEM co-location hourly-averaged results
Average bias
Average RMS before
Average RMS after

## Slide 17
Apply the FEM correction to the PA-II data from Charlotte, NC
Black shows monthly-averaged PM2.5 from BAM1022 (FEM)
Lighter purple is uncorrected PA-II PM2.5 (RMS )
Darker purple is the corrected PA-II PM2.5 (RMS )
20
15
10
5
Jan
2018
Jul
2018
Jan
2019
Jul
2019
Jul
2017

## Slide 18
Apply the FEM correction to the PA-II data from Winston-Salem, NC
Black shows monthly-averaged PM2.5 from T640X (FEM)
Lighter purple is uncorrected PA-II PM2.5 (RMS )
Darker purple is the corrected PA-II PM2.5 (RMS )
20
15
10
5
Jul
2018
Oct
2018
Jan
2019
Apr
2019
Apr
2018
Jul
2019

## Slide 19
Correction from Sensor B
Correction using Forsyth sensor
Uncorrected input PM2.5
Uncorrected input PM2.5
Uncorrected input PM2.5
Relative humidity (%)
Relative humidity (%)
Relative humidity (%)
Relative humidity (%)
Uncorrected input PM2.5
Sensor B vs. Sensor A
Sensitivity of correction to sensor site
Cool colors mean stronger correction
Cool colors mean stronger correction

## Slide 20
Charts show monthly averaged uncorrected PA-II minus FEM
Slight positive trends relative to FEM reference data
increase in bias per year
12
8
4
0
Jul
2018
Oct
2018
Jan
2019
Apr
2019
Apr
2018
Jul
2019
12
8
4
0
Jan
2018
Jul
2018
Jan
2019
Jul
2019
Jul
2017
Stability of Plantower sensors in PA-II for 15-31 months of sampling
Notes: Forsyth A (15) 0.15+0.12 (two glitch months in Apr-May 2018), Remount A (25) 0.07+0.04, Forsyth B (15) 0.09+0.12, Remount B (31) 0.06+0.03

## Slide 21
Note: Some of this work is published in our Aerosol Science and Technology paper
Future efforts
Continued expansion of Clean Air Carolina AirKeepers program in North Carolina
Case studies at higher PM2.5 concentrations
Large-scale synthesis of daily and hourly data
Health and exposure studies
Land-use regression towards a gridded data product